- Purpose:

To describe the methods for the determination of total moisture present in coal samples.

- Scope:

This procedure applies to the determination of total moisture in coal samples analyzed at SMJ Lab. It covers all coal samples submitted for routine testing, quality control, and reporting purposes.

- Definitions And Abbreviations:

Total Moisture: The coal which has been exposed to contact with water in the seam or in a washery, or coal and coke wetted by rain, may carry free or visible water. This water plus the moisture within the material, is referred to as total moisture.

- Safety Precautions:

- Use appropriate personal protective equipment (PPE), such as heat-resistant gloves, when handling hot crucibles.

- Ensure proper ventilation when operating the muffle furnace to prevent exposure to fumes.

- Follow all laboratory safety protocols and guidelines.

- Apparatus:

- Ventilated Air oven: Drying oven in which a constant and uniform temperature of 108 ±2°C can be maintained.

- Non corrodible or S.S Tray: Approximately 1000 sq.cm in area and 5 cm in depth.

- Weighing vessel: Shallow, approximately 40 cm² in area, made of silica, glass or stainless steel, with ground edges and fitted with ground cover.

- Reagents:

None required for this test method.

- Sample Handling and Preparation:

Obtain a special moisture sample of 1 kg of coal crushed to pass through a square mesh screen conforming to the designation 12.5 mm 'C' of IS: 460-1978, delivered in a sealed non-corrodible container.

- Procedure:

Determination of Total Moisture is carried out in two stages, i.e. Air Drying followed by Oven Drying.

  - Air Drying:

- Accurately weigh the sample and the container to the nearest 0.5 g (M₁).

- Transfer the sample from the container to a pre-weighed metal tray, spreading the material evenly on the tray.

- Dry the empty container in air and transfer any coal adhering to the sides onto the tray.

- Weigh the dry empty container (M₂).

- Calculate and record the mass of the sample taken for the test (M₁ - M₂).

- Allow the material on the tray to air-dry at atmospheric temperature in a well-ventilated, dust-free area.

- Periodically weigh the tray and sample during the air-drying process.

- Consider the air-drying complete when the change in mass during an hour is less than 0.1% of the sample mass.

- Record the final mass after air-drying (M₃).

  - Oven Drying:

- Heat an empty weighing vessel in a drying oven at 108 ± 2°C.

- After heating, allow the empty vessel to cool for 20 minutes, and then weigh it (W₁).

- Crush the air-dried coal sample to pass through a 2.90mm IS sieve.

- Spread approximately 10 g of the crushed material uniformly in the pre-weighed vessel and weigh it (W₂).

- Place the uncovered vessel containing the sample in the drying oven and heat it at a temperature of 108 ±2°C.

- Continue heating until there is no further loss in mass, which typically takes 1.5 to 3 hours.

- After drying, replace the cover on the vessel and transfer it to a desiccator.

- Allow the covered vessel to cool in the desiccator for 20 minutes.

- After cooling, weigh the covered vessel containing the dried sample (W₃).

  - Calculation:

- The percentage loss in mass of the original sample in air-drying as:

(Air Dry Loss, ADL), X =100 × (M2 – M3 )(M2 – M1)

- The percentage loss in mass of the air-dried sample on oven drying as:

Residual Moisture %, Y = 100 × (W2-W3)(W2-W1)

Total Moisture % = X + Y 1-X100

- Reference:

- Indian Standard IS 436 (Part 1): 2024, Methods for Sampling of Coal and Coke.

- Indian Standard IS 1350 (Part 1): 2025, Methods of Test for Coal and Coke - Proximate Analysis.