- Purpose:

To describe the method for the determination of total ash content in coal samples.

- Scope:

This procedure applies to the determination of total ash content in coal samples received and analyzed by SMJ Lab. It covers all coal samples submitted for routine testing, quality control, and reporting purposes.

- Definitions And Abbreviations:

Total Ash Content: The residue remaining after complete combustion of the coal sample.

- Safety Precautions:

- Use appropriate personal protective equipment (PPE), such as heat-resistant gloves, when handling hot crucibles.

- Ensure proper ventilation when operating the muffle furnace to prevent exposure to fumes.

- Follow all laboratory safety protocols and guidelines.

- Apparatus:

- Muffle furnace capable of giving a substantially uniform zone of 500°C in 30 minutes from cold, of being raised to 815 ± 10°C in a further 30 to 60 minutes and of maintaining this temperature up to the end of the run-up period. The furnace should also be capable of being raised to a temperature of 850 ± 10°C, if necessary, the ventilation shall be such as to give about five air changes per minute.

- Dish made of silica, porcelain or platinum, 10 to 15 mm deep, of such a size that, with the amount of sample used, the spread does not exceed 0·15 g/cm²

- Analytical balance capable of weighing up to 0.1 mg.

- Desiccator.

- Tongs or forceps for handling hot crucibles.

- Reagents:

None required for this test method.

- Sample Handling and Preparation:

Pulverized coal sample is air-dried and passed through a 212-micron IS sieve.

- Procedure:

- Mix the air-dried material thoroughly using a laboratory mechanical mixer for one minute.

- Weigh a clean dry empty dish and its lid accurately (W₁).

- Add about one to two grams of the material into the dish, ensuring even distribution with a spread not exceeding 0.15 g per cm² (W₂).

- Place the uncovered dish into the muffle furnace at room temperature.

- Increase the furnace temperature to 500°C, over 30 minutes.

- Further raise the temperature to 815 ± 10°C in 30 to 60 minutes and maintain it for 60 minutes.

- Cover the dish with its lid if the ash is light and fluffy, then remove it from the furnace and allow it to cool. First, cool it on a cold metal slab for ten minutes and then transfer it to a desiccator beside the balance.

- Weigh the dish after it has been in the desiccator for 15 minutes (W₃).

- Re-ignite the dish containing ash at the same temperature until the change in mass of ash is less than 0.001 g.

- Note the colour of the ash after re-ignition and brush out the ash carefully before reweighing the empty dish to obtain the mass of the ash by the difference in weights (W₄).

- Calculation:

The total ash % can be calculated by the below formula:

Total Ash % = 100 × (W3 – W4 )(W2 – W1)

Where,

W₁ = mass of empty dish in g.

W₂ = mass of dish + sample in g.

W₃ = mass of dish + ash in g.

W₄ = mass of dish after brushing out the ash and then reweighing

- Reference:

- Indian Standard IS 436 (Part 1): 2024, Methods for Sampling of Coal and Coke.

- Indian Standard IS 1350 (Part 1): 2025, Methods of Test for Coal and Coke - Proximate Analysis.