- Purpose:

To define the procedure for determination of Gross Calorific Value of coal by using a bomb calorimeter.

- Scope:

This procedure applies to all coal samples received for Gross Calorific Value (GCV) determination, including routine testing, method verification, and proficiency testing activities, and shall be performed only by laboratory personnel authorized for calorimetric analysis.

- Definitions And Abbreviations:

  - GCV (Gross Calorific Value)

The total amount of heat released when a specified quantity of coal is completely combusted in oxygen under standard conditions, including the latent heat of condensation of water vapour formed during combustion

  - Water Equivalent (W)

The effective heat capacity of the calorimeter system, expressed as the amount of water that would absorb the same quantity of heat for a unit rise in temperature.

  - ARB (As Received Basis)

Result expressed on the sample in the condition as received.

  - DB (Dry Basis)

Result expressed after eliminating total moisture content.

- Safety Precautions:

- Use appropriate personal protective equipment (PPE), such as heat-resistant gloves, when handling hot crucibles.

- Ensure proper ventilation when operating the muffle furnace to prevent exposure to fumes.

- Follow all laboratory safety protocols and guidelines.

- Apparatus:

- Bomb Calorimeter (Manual/Automatic)

- Oxygen cylinder (minimum 99.5% purity)

- Analytical balance (0.0001 g readability)

- Pellet press (if required)

- Ignition wire (Nickel/Chromium or equivalent)

- Cotton thread (if used)

- Thermometer / digital temperature sensor (0.001°C resolution recommended)

- Reagents:

- Distilled / deionized water

- Certified Reference Material (Benzoic Acid, known calorific value)

- Oxygen (≥99.5% purity)

- Sodium carbonate solution (for washing bomb)

- Sample Handling and Preparation:

- Sample shall be representative and air-dried as required.

- Grind to pass through 212-micron sieve.

- Mix thoroughly before subsampling.

- To be stored in airtight container before analysis

Moisture determination should be conducted separately if results are to be reported on dry basis.

- Procedure:

  - Calibration of Calorimeter

Calibration shall be performed before first use, after maintenance, when the control chart shows drift, and as per the internal calibration schedule.

Calibration Procedure

- Weigh certified benzoic acid (~1 g).

- Combust under standard conditions.

- Record temperature rise.

- Calculate Energy Equivalent (W) of calorimeter:

W= m×CVCRM+eΔT

Where:

m = mass of CRM (g)

CVCRM = certified calorific value (cal/g or J/g)

e = corrections (wire, thread, acid)

ΔT = corrected temperature rise

Maintain control chart of energy equivalent.

  - Test Procedure

- Preparation

- Weigh approximately 1 g of prepared coal sample (m).

- Record weight to 0.0001 g.

- Attach ignition wire.

- Add 10 ml distilled water inside bomb.

- Charging

- Seal bomb properly.

- Fill with oxygen to 25–30 bar (as per instrument specification).

- Combustion

- Place bomb in calorimeter vessel containing measured water quantity.

- Allow temperature stabilization.

- Initiate ignition.

- Record temperature rise until maximum stable temperature is achieved.

- Corrections

Apply corrections for:

- Ignition wire combustion

- Cotton thread

- Acid formation (if applicable)

- Calculation:

The total GCV can be calculated by the below formula:

Gross Calorific Value as recived= W×ΔT-em

Where:

W = Energy equivalent of calorimeter

ΔT = Corrected temperature rise

e = Total corrections

m = Mass of sample (g)

Gross Calorific Value Dry Basis= Ash(AR)×100100-Moisture%

- Reference:

- Indian Standard IS 1350 (Part 2) : 2022, Method of Test for Coal and Coke — Determination of Gross Calorific Value.