- Purpose:

To define the procedure for determination of ash content in solid biofuels by complete combustion in a muffle furnace.

- Scope:

This procedure is applicable to the analysis of Biomass pellets, Briquettes, Agro-residues, Wood chips, Torrefied biomass and other solid biofuels.

- Definitions And Abbreviations:

  - Biomass Pellets

Densified solid biofuel produced by mechanically compressing milled biomass (e.g., sawdust, agro-residues) into cylindrical shapes, typically without chemical binders.

  - Briquettes

Solid biofuel blocks formed by compressing biomass under high pressure, generally larger than pellets and may contain natural binders inherent in biomass.

  - Agro-residues

Agricultural by-products remaining after harvesting or processing crops (e.g., rice husk, bagasse, groundnut shell, mustard stalk).

  - Wood Chips

Mechanically reduced wood pieces produced by chipping logs, branches, or wood waste, used as fuel without densification.

  - Torrefied Biomass

Biomass that has undergone mild thermal treatment (typically 200–300°C in absence of oxygen), resulting in improved fuel properties such as higher energy density and hydrophobicity.

  - Solid Biofuels

Solid fuels produced directly or indirectly from biomass, including but not limited to pellets, briquettes, chips, agricultural residues, and thermally treated biomass.

- Safety Precautions:

- Use appropriate personal protective equipment (PPE), such as heat-resistant gloves, when handling hot crucibles.

- Ensure proper ventilation when operating the muffle furnace to prevent exposure to fumes.

- Follow all laboratory safety protocols and guidelines.

- Apparatus:

- Muffle Furnace (Temperature range up to 600°C)

- Analytical Balance (0.0001 g readability)

- Porcelain / Silica crucibles

- Desiccator with fresh desiccant

- Tongs

- Hot air oven (for pre-drying if required)

- Stopwatch / Timer

- Reagents:

None required for this test method.

Desiccant (Silica gel / equivalent).

- Sample Handling and Preparation:

- Sample shall be prepared as per relevant sample preparation standard.

- Particle size: Typically, ≤1 mm (as applicable to product standard).

- Sample must be representative and properly mixed.

If moisture determination is required separately, perform before ash determination.

- Procedure:

  - Preparation of Crucible

- Clean crucible.

- Heat empty crucible in muffle furnace at 550°C for 30 minutes.

- Cool in desiccator to room temperature.

- Weigh and record mass (W₁).

  - Weighing of Sample

- Weigh approximately 1.0 g of prepared sample into crucible.

- Record combined weight (W₂).

- Spread sample evenly.

  - Ashing

- Place crucible in cold furnace.

- Raise temperature gradually to 250°C and hold for 30 minutes (to avoid flaming losses).

- Increase temperature to 550 ± 10°C.

- Maintain at this temperature for minimum 2 hours.

- Ensure complete combustion (residue grey/white).

If black particles remain:

- Cool

- Moisten residue with few drops distilled water

- Dry

- Re-ignite for 30 minutes

  - Cooling and Weighing

- Remove crucible carefully using tongs.

- Cool in desiccator to room temperature (approx. 30–45 min).

- Weigh and record final mass (W₃).

- Calculation:

The total ash % can be calculated by the below formula:

Ash % as recived= W3-W1W2-W1×100

Where:

W₁ = Mass of empty crucible (g)

W₂ = Mass of crucible + sample before ignition (g)

W₃ = Mass of crucible + ash (g)

Ash % Dry Basis= Ash(AR)×100100-Moisture%

- Reference:

- Indian Standard IS 17653: 2021, Solid Biofuels — Determination of Ash Content.