- Purpose:

To define the procedure for determination of Gross Calorific Value of solid biofuels by using a bomb calorimeter.

- Scope:

This procedure is applicable to the analysis of Biomass pellets, Briquettes, Agro-residues, Wood chips, Torrefied biomass and other solid biofuels.

- Definitions And Abbreviations:

  - Biomass Pellets

Densified solid biofuel produced by mechanically compressing milled biomass (e.g., sawdust, agro-residues) into cylindrical shapes, typically without chemical binders.

  - Briquettes

Solid biofuel blocks formed by compressing biomass under high pressure, generally larger than pellets and may contain natural binders inherent in biomass.

  - Agro-residues

Agricultural by-products remaining after harvesting or processing crops (e.g., rice husk, bagasse, groundnut shell, mustard stalk).

  - Wood Chips

Mechanically reduced wood pieces produced by chipping logs, branches, or wood waste, used as fuel without densification.

  - Torrefied Biomass

Biomass that has undergone mild thermal treatment (typically 200–300°C in absence of oxygen), resulting in improved fuel properties such as higher energy density and hydrophobicity.

  - Solid Biofuels

Solid fuels produced directly or indirectly from biomass, including but not limited to pellets, briquettes, chips, agricultural residues, and thermally treated biomass.

- Safety Precautions:

- Use appropriate personal protective equipment (PPE), such as heat-resistant gloves, when handling hot crucibles.

- Ensure proper ventilation when operating the muffle furnace to prevent exposure to fumes.

- Follow all laboratory safety protocols and guidelines.

- Apparatus:

- Bomb Calorimeter (Manual/Automatic)

- Oxygen cylinder (minimum 99.5% purity)

- Analytical balance (0.0001 g readability)

- Pellet press (if required)

- Ignition wire (Nickel/Chromium or equivalent)

- Cotton thread (if used)

- Distilled water

- Thermometer / digital temperature sensor (0.001°C resolution recommended)

- Reagents:

- Distilled / deionized water

- Certified Reference Material (Benzoic Acid, known calorific value)

- Oxygen (≥99.5% purity)

- Sample Handling and Preparation:

- Sample shall be representative and air-dried as required.

- Grind to appropriate particle size (typically ≤1 mm).

- Mix thoroughly before subsampling.

- Press into pellet if required to ensure complete combustion.

Moisture determination should be conducted separately if results are to be reported on dry basis.

- Procedure:

  - Calibration of Calorimeter

Calibration shall be performed before first use, after maintenance, when the control chart shows drift, and as per the internal calibration schedule.

Calibration Procedure

- Weigh certified benzoic acid (~1 g).

- Combust under standard conditions.

- Record temperature rise.

- Calculate Energy Equivalent (W) of calorimeter:

W= m×CVCRM+eΔT

Where:

m = mass of CRM (g)

CVCRM = certified calorific value (cal/g or J/g)

e = corrections (wire, thread, acid)

ΔT = corrected temperature rise

Maintain control chart of energy equivalent.

  - Test Procedure

- Preparation

- Weigh approximately 1 g of prepared sample (m).

- Record weight to 0.0001 g.

- Attach ignition wire.

- Add 10 ml distilled water inside bomb (if required as per instrument SOP).

- Charging

- Seal bomb properly.

- Fill with oxygen to 25–30 bar (as per instrument specification).

- Combustion

- Place bomb in calorimeter vessel containing measured water quantity.

- Allow temperature stabilization.

- Initiate ignition.

- Record temperature rise until maximum stable temperature is achieved.

- Corrections

Apply corrections for:

- Ignition wire combustion

- Cotton thread

- Nitric acid formation

- Sulfuric acid formation (if applicable)

- Calculation:

The total GCV can be calculated by the below formula:

Gross Calorific Value as recived= W×ΔT-em

Where:

W = Energy equivalent of calorimeter

ΔT = Corrected temperature rise

e = Total corrections

m = Mass of sample (g)

Gross Calorific Value Dry Basis= Ash(AR)×100100-Moisture%

- Reference:

- Indian Standard IS 17654: 2021, Solid Biofuels — Determination of Calorific Value.