# Purpose

To describe the method for determination of total ash content in coal samples.

# Scope

This procedure describes the method for determination of total ash content in coal samples received and analyzed by SMJ Laboratory. The procedure applies to samples submitted for routine testing, quality control, and reporting purposes.

# Normative Reference

## Indian Standard IS 436 (Part 1): 2024

Methods for Sampling of Coal and Coke.

## Indian Standard IS 1350 (Part 1): 2025

Methods of Test for Coal and Coke - Proximate Analysis.

# Definitions And Abbreviations

## Total Ash Content

The incombustible residue remaining after complete combustion of coal under specified conditions.

# Safety Precautions

- Wear heat-resistant gloves and safety goggles when handling hot crucibles or dishes.

- Use tongs or a crucible holder to remove dishes from the furnace.

- Ensure adequate ventilation during furnace operation.

- Follow all laboratory safety procedures and PPE requirements.

# Apparatus

## Muffle furnace

Capable of being raised to a temperature of 850 ± 10°C, if necessary, the ventilation shall be such as to give about five air changes per minute.

## Dish made of silica, porcelain or platinum

10 to 15 mm deep, of such a size that, with the amount of sample used, the spread does not exceed 0·15 g/cm².

## Analytical balance

Capable of weighing up to 0.1 mg.

## Desiccator

## Tongs or crucible holder

For handling hot crucibles.

# Reagents

None required for this test method.

# Sample Handling and Preparation

Pulverize the coal sample and air-dry it. Pass the sample through a 212 μm IS sieve and mix thoroughly before analysis.

# Procedure

- Mix the air-dried material thoroughly using a laboratory mechanical mixer for one minute.

- Weigh a clean dry empty dish and its lid accurately (W₁).

- Add 1–2 g of sample to the dish ensuring uniform distribution with a loading not exceeding 0.15 g/cm² (W₂).

- Place the uncovered dish into the muffle furnace at room temperature.

- Increase the furnace temperature to 500°C, over 30 minutes.

- Further raise the temperature to 815 ± 10°C in 30 to 60 minutes and maintain it for 60 minutes.

- Cover the dish with its lid if the ash is light and fluffy, then remove it from the furnace and allow it to cool. First, cool it on a cold metal slab for ten minutes and then transfer it to a desiccator beside the balance.

- Weigh the dish after it has been in the desiccator for 15 minutes (W₃).

- Re-ignite the dish containing ash at the same temperature until the change in mass of ash is less than 0.001 g.

- Note the color of the ash after re-ignition and brush out the ash carefully before reweighing the empty dish to obtain the mass of the ash by the difference in weights (W₄).

# Calculation

The total ash % can be calculated by the below formula:

| Total Ash % = | (W₃ – W₄) | x 100 |
| --- | --- | --- |
|  | (W₂ – W₁) |  |