# Standard Operating Procedure

# for

# Operation and Intermediate Checks – Oven,  Furnace,  Humidity Chamber

# Purpose:

This SOP outlines the procedure for the safe and efficient operation of Ovens, Furnaces, or Humidity Chambers in the laboratory environment. It also details the necessary intermediate checks to maintain equipment functionality and reliability.

# Scope:

This SOP applies to all laboratory personnel responsible for operating Ovens, Furnaces, or Humidity Chambers in SMJSF Lab.

# Responsibilities:

- Laboratory Supervisor: Responsible for overseeing the implementation of this SOP and ensuring compliance with safety regulations.

- Equipment Operator: Responsible for following the procedures outlined in this SOP and conducting intermediate checks as required.

# Equipment Description:

- Ovens, Furnaces, or Humidity Chambers are used for various purposes such as drying, heating, sterilizing, or maintaining controlled environmental conditions.

- Each equipment unit is equipped with temperature control settings, timers, and safety features.

# Personal Protective Equipment (PPE):

- Heat-Resistant Gloves

- Safety Goggles or Glasses

- Lab Coat or Protective Clothing

- Closed-Toe Shoes

- Long Pants

- Respiratory Protection (if applicable)

- Hair Protection

- Additional PPE as Required

# Procedure:

## Pre-Operation Checks:

- Before starting the equipment, conduct a thorough inspection to ensure it is clean and free from any debris, residue, or obstruction. Any accumulated material inside the equipment can affect its performance and may pose safety hazards.

- Verify that the power supply cord is properly connected to a grounded outlet and that there are no visible signs of damage to the cord or plug. Ensure that the power switch is in the off position before proceeding.

- Check the exterior of the equipment for any signs of wear, damage, or corrosion. Pay close attention to the door seals, hinges, and control panels. Any damaged components should be reported to maintenance for repair or replacement.

- Inspect the ventilation system of the equipment to ensure proper airflow. Remove any blockages or obstructions from vents or exhaust ports to prevent overheating and ensure efficient operation.

- Review the equipment manual or operating instructions provided by the manufacturer for specific pre-operation checks and safety precautions. Familiarize yourself with the equipment's features, controls, and operating procedures before proceeding.

## Start-Up Procedure:

- Turn on the power switch located on the control panel of the equipment. Some models may have additional power switches or buttons located elsewhere, so refer to the equipment manual for guidance.

- Allow the equipment to initialize and perform any self-diagnostic tests that may be built into the system. This may take a few moments, depending on the model and complexity of the equipment.

- Once the equipment is powered on, set the desired operating parameters such as temperature, humidity, and timer settings using the control panel or interface. Follow the recommended settings based on the type of samples or materials being processed.

- Ensure that the temperature and other parameters are set within the safe operating range of the equipment to avoid overheating or overloading. Consult the equipment manual for guidance on appropriate settings for different applications.

## Operation:

- Load the samples or materials into the equipment according to the specified protocols and procedures. Use appropriate containers or trays to hold the samples securely and prevent spillage or contamination.

- Close the door or lid of the equipment securely to maintain the desired temperature, humidity, or environmental conditions inside the chamber. Ensure that the door seals are intact and free from any damage or gaps.

- Monitor the equipment periodically during operation to ensure that it is functioning properly and maintaining the desired parameters. Check the temperature, humidity, or other relevant indicators displayed on the control panel to verify stability and consistency.

- Avoid opening the door or lid of the equipment unnecessarily during operation, as this can lead to temperature fluctuations and affect the quality of the results. If access to the interior of the equipment is required, minimize the duration of the door being open and close it promptly afterward.

## Intermediate Checks:

- Conduct intermediate checks during operation to verify the accuracy and stability of the equipment's performance. This may include checking the temperature using a calibrated thermometer or verifying the humidity levels using a hygrometer.

- Monitor the equipment for any abnormal noises, vibrations, or odors that may indicate a malfunction or potential issue. If any abnormalities are observed, stop the equipment immediately and investigate the cause before resuming operation.

- Inspect the condition of the heating elements, seals, and internal components for any signs of wear, damage, or deterioration. Clean or replace any components as necessary to maintain optimal performance.

- Check the cleanliness of the equipment interior and remove any accumulated debris, residue, or contaminants. Use appropriate cleaning agents and methods recommended by the manufacturer to avoid damage to the equipment.

## Shut-Down Procedure:

- Once the operation is complete or when the equipment is no longer needed, follow the shut-down procedure to power off the equipment safely.

- Turn off the power switch located on the control panel and wait for the equipment to cool down to a safe temperature before proceeding.

- Remove any samples or materials from the equipment and clean the interior surfaces if necessary. Use mild detergents or solvents to remove stubborn residues, taking care not to damage sensitive components.

- Close the door or lid of the equipment securely to prevent dust, debris, or contaminants from entering the chamber when not in use.

- Disconnect the power supply cord from the outlet and ensure that all controls are in the off position before leaving the equipment unattended.