INTERNAL QUALITY CONTROL (IQC) PLAN

SMJ Sustainable Fuel Lab Pvt. Ltd.

1. Purpose

To ensure the ongoing validity, accuracy, precision, and reliability of test results generated for SMJ lab, in compliance with ISO/IEC 17025:2017 Clause 7.7 and NABL requirements.

2. Scope of IQC

This IQC plan applies to the following materials, parameters, and standards:

2.1 Biomass Briquettes

| Parameter | Standard |
| --- | --- |
| Ash Content | IS 17653 : 2021 |
| Gross Calorific Value (GCV) | IS 17654 : 2021 |
| Moisture Content | IS 17655 (Part 2) : 2021 |

2.2 Coal

| Parameter | Standard |
| --- | --- |
| Ash Content | IS 1350 (Part 1) : 2025 |
| Moisture Content | IS 1350 (Part 1) : 2025 |
| Volatile Matter | IS 1350 (Part 1) : 2025 |
| Fixed Carbon | IS 1350 (Part 1) : 2025 (Calculated by difference) |
| Gross Calorific Value (GCV) | IS 1350 (Part 2) : 2022 |

3. Responsibility

| Role | Responsibility |
| --- | --- |
| Technical Manager | Approval of IQC plan, technical review |
| Quality Manager | Monitoring IQC data, corrective actions |
| Authorized Analyst | Performing IQC activities & recording |

4. Regular Use of Certified Reference Materials / Quality Control Materials

4.1 Type of IQC Materials

- Certified Reference Materials (CRM) for GCV (Benzoic Acid or equivalent)

- In-house Secondary Reference Materials (SRM) for:

  - Ash

  - Moisture

  - Volatile Matter

- Homogeneous, stable retained coal and biomass briquette samples

4.2 Frequency

- At least once per test batch or per day, whichever is earlier

- Mandatory after:

  - Calibration / maintenance

  - Instrument breakdown

  - Change of analyst

4.3 Acceptance Criteria

- CRM: Certified value ± stated uncertainty

- SRM / Working standard: Mean ± 2 SD (warning), ± 3 SD (action)

5. Use of Alternative Calibrated Instrumentation (Traceability)

All instruments used are:

- Externally calibrated by NABL-accredited calibration laboratories

- Traceable to national / international standards

| Instrument | Model / ID | Calibration Validity |
| --- | --- | --- |
| Digital Bomb Calorimeter | DBCM100 | 02-01-2026 |
| Single Pan Balance | DWB-01 | 15-10-2025 |
| Muffle Furnace | BTI-36 | 11-01-2025 |
| Hot Air Oven | O-3427 | 17-10-2025 |

Alternate instruments (if any) are used only after performance verification with IQC samples.

6. Functional Checks of Measuring & Testing Equipment

6.1 Routine Checks

| Equipment | Functional Check | Frequency |
| --- | --- | --- |
| Balance | Zero, repeatability | Daily |
| Muffle Furnace | Temperature set-point verification | Daily |
| Hot Air Oven | Temperature stability | Daily |
| Bomb Calorimeter | Fuse wire, pressure integrity | Daily |

Records maintained in equipment logbooks.

7. Use of Working Standards with Control Charts

7.1 Parameters Covered

Control charts are maintained for:

- Ash

- Moisture

- Volatile Matter

- GCV

Fixed Carbon is monitored indirectly via mass balance consistency.

7.2 Control Limits

- Warning limits: ± 2 SD

- Action limits: ± 3 SD

7.3 Review

- Daily by analyst

- Monthly by Quality Manager

8. Intermediate Checks on Measuring Equipment

| Equipment | Intermediate Check | Frequency |
| --- | --- | --- |
| Balance | Certified check weights | Monthly |
| Furnace / Oven | Secondary calibrated thermometer | Quarterly |
| Bomb Calorimeter | Benzoic acid check | Weekly |

9. Replicate Tests / Repeatability Checks

- Duplicate testing carried out for:

  - New sample types

  - Customer complaints

  - IQC failures

- Acceptable limits as per relevant IS method

10. Resetting / Recalibration of Retained Items

Recalibration or reset is initiated when:

- IQC results exceed ±3 SD

- Persistent trend observed

- After major repair or maintenance

Retained IQC samples are revalidated annually.

11. Correlation of Results (Internal Consistency)

11.1 Coal Proximate Analysis Check

Moisture + Ash + VM + Fixed Carbon = 100 ± 1 %

11.2 GCV Correlation

- GCV results reviewed against:

  - Ash percentage

  - Historical coal / briquette data

Deviations trigger investigation.

12. Review of Reported Results

Before release, results are reviewed for:

- IQC compliance

- Instrument calibration validity

- Control chart status

Only results meeting IQC criteria are authorized.

13. Inter-Laboratory Comparisons (ILC / PT)

- Participation in NABL-recognized PT schemes

- Minimum once per year per parameter

- Performance evaluated using Z-score

Corrective actions initiated for unsatisfactory performance.

14. Testing of Blind Samples

- Blind samples introduced periodically by Quality Manager

- Analyst unaware of identity

- Results evaluated against assigned values

- Records maintained and reviewed

15. Handling of Out-of-Control Situations

- Immediate suspension of affected testing

- Root cause analysis

- Impact assessment on reported results

- Corrective and preventive action

16. Records & Retention

| Record | Retention |
| --- | --- |
| IQC data & control charts | 5 years |
| Equipment logs | 5 years |
| PT / ILC records | 1 NABL cycle |

17. Review & Continuous Improvement

- Monthly IQC review by Quality Manager

- Annual review during Management Review

- IQC plan updated with scope or method changes