## GCV
Method Verification Report SMJ Sustainable Fuel Lab Pvt. Ltd.
| 1. | Objective
To verify the performance characteristics (repeatability, accuracy and measurement uncertainty) of the bomb calorimeter for determination of Gross Calorific Value (GCV), in accordance with ISO/IEC 17025 requirements.
| 2. | Details
|  | Report No. |  |  |  |  | : | SMJ26MVR001
|  | Standard Method |  |  |  |  | : | IS 1350 (Part 2) : 2022
|  | Number of Replicates |  |  |  |  | : | 6
|  | Instrument / Equipment |  |  |  |  | : | a) Digital Bomb Calorimeter, 0.01 °C
b) Single pan balance, 0.0001 g
|  | Reference Material |  |  |  |  | : | Coal CRM, GCV value: 5200 cal/g, uncertainty: 50
| 3. | Experimental Data
|  | Water Equivalent (W̄) : |  |  |  | 6318 cal/°C
|  | Run | Sample wt., g | [5 empty] | Temp. rise, °C |  |  | Measured GCV, cal/g |  |  |  | Mean GCV |  | Std. Dev. (SD)
|  | 1 | 1.0045 | [5 empty] | 0.831 |  |  | 5204.44 |  |  |  | 5212.67 |  | 43.664
|  | 2 | 1.0011 | [5 empty] | 0.821 |  |  | 5159
|  | 3 | 1.0047 | [5 empty] | 0.83 |  |  | 5197.11
|  | 4 | 0.9987 | [5 empty] | 0.838 |  |  | 5278.95
|  | 5 | 1.0029 | [5 empty] | 0.827 |  |  | 5187.54
|  | 6 | 1.0032 | [5 empty] | 0.837 |  |  | 5248.97
| 4. | Method Verification
[10 empty] | Acceptance Criteria |  |  | Calculated Value |  |  |  | Status
|  | Precision (Repeatability), %RDS | [7 empty] | ≤ 2% |  |  | 0.838 % |  |  |  | Fail
|  | Accuracy (Bias Study), BIAS | [7 empty] | 50 |  |  | 12.67 |  |  |  | Pass
| 5. | Verification Statement
Method verification failed due to unacceptable repeatability. Corrective action initiated.
| 6. | Ownership & Authorization
|  | Prepared By |  | : | Laboratory Head | [5 empty] | Reviewed By |  |  | : | Quality Manager
|  | Signature |  | : | [6 empty] | Signature |  |  | :
|  | Date |  | : | 2026-02-17 | [5 empty] | Reviewed By |  |  | : | Technical Manager
|  | Seal |  | : | [6 empty] | Signature |  |  | :
Document No.: SMJ/FRM/QCD/07 | [10 empty] | Page 1 of 1

## Proximate
Method Verification Report SMJ Sustainable Fuel Lab Pvt. Ltd.
| 1. | Objective
To verify the performance characteristics (repeatability, accuracy and measurement uncertainty) of the bomb calorimeter for determination of Gross Calorific Value (GCV), in accordance with ISO/IEC 17025 requirements.
| 2. | Details
|  | Report No. |  |  |  |  | : | SMJ26MVR002
|  | Standard Method |  |  |  |  | : | IS 1350 (Part 1) : 2025
|  | Number of Replicates |  |  |  |  | : | 6
|  | Instrument / Equipment |  |  |  |  | : | a) Analytical Balance (0.0001 g)
b) Hot Air Oven (105–110°C)
c) Muffle Furnace (900 ±15°C)
d) Desiccator
|  | Reference Material |  |  |  |  | : | Coal CRM, Moisture value: 42.1%, uncertainty: 0.45
Coal CRM, Ash Content value: 14%, uncertainty: 0.2
Coal CRM, Volatile Matter value: 41%, uncertainty: 0.5
| 3. | Experimental Data
|  | Run | Moisture (%) | [5 empty] | Ash Content (%) |  |  |  | Volatile Matter (%) |  |  |  |  | Fixed Carbon (%)
1
2
3
4
5
6
Statistical Summary
|  | Parameter |  |  |  |  | Mean (%) |  |  | SD |  | %RSD |  |  |  |  | Certified Value |  |  | % Bias
Moisture
Ash Content
Volatile Matter
Fixed Carbon
| 4. | Method Verification
|  | Parameter |  |  |  | Verification Type |  |  |  | Acceptance Criteria |  |  |  | Calculated Value |  |  |  |  | Status
|  | Moisture |  |  |  | Precision (%RSD) |  |  |  | ≤ 2%
[6 empty] | Accuracy (% Bias) |  |  |  | ±5%
|  | Ash Content |  |  |  | Precision (%RSD) |  |  |  | ≤ 2%
[6 empty] | Accuracy (% Bias) |  |  |  | ±5%
|  | Volatile Matter |  |  |  | Precision (%RSD) |  |  |  | ≤ 2%
[6 empty] | Accuracy (% Bias) |  |  |  | ±5%
|  | Fixed Carbon* |  |  |  | Precision (%RSD) |  |  |  | ≤ 2%
*Fixed Carbon is calculated by difference.
| 5. | Verification Statement
Method verification failed due to unacceptable repeatability. Corrective action to be initiated.
| 6. | Ownership & Authorization
|  | Prepared By |  | : | Laboratory Head | [5 empty] | Reviewed By |  |  |  | : | Quality Manager
|  | Signature |  | : | [6 empty] | Signature |  |  |  | :
|  | Date |  | : | 2026-02-17 | [5 empty] | Reviewed By |  |  |  | : | Technical Manager
|  | Seal |  | : | [6 empty] | Signature |  |  |  | :
Document No.: SMJ/FRM/QCD/07 | [10 empty] | Page 1 of 1